## Slide 1: OBS-ONE Platform Gaseous, PM and PN
Horiba Instruments, Inc.

## Slide 2
System Features
OBS-ONE-PM (Particulate Mass)
OBS-ONE-PN (Particle Number)

## Slide 3
Gaseous Components:
CO, CO2, NO, NOx, THC, CH4
Particulate:
Mass, Solid Particle Number
OBS-ONE -PM
OBS-ONE -PN

## Slide 4: OBS ONE Concept
Same base software package for new generation laboratory equipment (MEXA ONE, CVS ONE, etc.)
  Horiba One Platform
  Windows based interface
  Linux machine for equipment control
Software includes test automation (sequence for PEMS)
  Flexible test sequences
  Large amount of entry fields for documentation of test
Post processing software based on National Instruments DIADEM software
  Standard graphical and mathematics package
  Emission calculation and analysis
  Default basic reports
  Customizable report formats

## Slide 5
Example of Type approval result (MAW)
Post data analysis(1)
Notes: 5

## Slide 6
Optional function with video, map data
Example of data analysis
Post data analysis(2)
Notes: 6

## Slide 7: System Flexibility..
GDC-ONE
GDC-703
Maintenance
  Gas analyzers use many common parts with MEXA-ONE, MEXA-7000 and OBS-2000.
  Linearization sequence identical to MEXA ONE
    Automated sequence with Horiba MFC based gas dividers
    User can update curve coefficients
  Automated interference checks with heated bubbler
    User can update correction coefficients
Notes:     Consumable parts :
      Sample gas filter element, NO2 to NO converter material etc
    Periodical replacement parts :
      Ozone generator, Sampling pump, electrical circuit bards etc
  Software for controller is designed with HORIBA ONE PLATFORM.
    MEXA-ONE quality check, maintenance support function (IO controller with flow diagram etc) and notification function.
    The software can control MEXA-ONE device for maintenance.
    For example : GDC-703 and GDC-ONE for linearity check and NOx converter efficiency check

## Slide 8
“Wet” based concentration measurement for better performance
  Direct conversion to mass from exhaust flow without dry to wet correction
  Faster response using “Hot wet measurement” technique (no cooler)
  More accurate real CO and CO2 measurement by Heated NDIR
  Reliable / accurate NOx measurement by CLD
Component | Detection principle | Advantage
NO, NOx | Heated Dual-CLD | No dehumidifier : fast response Small size, low sample flow rate
NO2 | (Calculated by Dual-CLD) | No dehumidifier so no loss of NO2
CO | Heated NDIR | No dehumidifier : fast response Integral H2O measurement for compensation
CO2 |  | 
THC | Heated FID | Vacuum type sampling for fast response. 191̊ C for diesel THC measurement
CH4 | Heated FID with non-methane cutter | Vacuum type sampling for fast response.
High performance detection
Notes: 8

## Slide 9: Feature of Heated-NDIR
FTP-75 Phase 1 (Cold start) Lean burn gasoline vehicle
Real concentration at measuring point can be measured by H-NDIR
NDIR with dehumidifier increases error due to dry to wet base correction during cold start of engine or for dual fuelled vehicles
【Patent No. JP 03771849】
Velocity
[km/h]
CO2
[vol%]
Heated-NDIR (H-NDIR)
CO and CO2 signal are compensated by H2O signal and optimized algorithm for more accurate measureament

## Slide 10: Sampling optimized
P
191℃
Heated region
Heated Pump
Vacuum type sampling system for size reduction
FILTER
Sample in
Detector
    Heated pump is bigger than cold pump
    More electrical power is required
P
191℃
Non Heated Pump
FILTER
P
Heated region
Detector
Sample in
    Smaller size and lower power consumption of pump
    Less heat insulation required
    Smaller instrument size and smaller battery due to lower power consumption => Smaller size of total system

## Slide 11
Work flow Test cell & PEMS
Test cell & PEMS
Testing vehicle with PEMS,
Compact test automation system
Data analysis & reporting
Store the data on DB
Simulate real world data on test cell

## Slide 12
System Features
OBS-ONE-PM (Particulate Mass)
OBS-ONE-PN (Particle Number)

## Slide 13
Key points
  Same size as OBS-ONE-GAS
  Keep the same concept: Front Access, no tools needed for operation
FS: Filter Sampler
RP: Real-time PM Sensor
FA
GA
CC
PM analyzer
Gas analyzer
470 mm
350 mm
350 mm
460 mm
Design/concept
OBS-ONE-PM

## Slide 14: PM Sampler and Real-Time Sensor
Exhaust pipe
P
mixer
47℃
Dilution Air
○ Correlation
○ Traceability
○ High sensitivity
○ Real time measurement
Diffusion Charger Sensor (DCS)
Exhaust flow rate signal from pitot tube flow meter used to control proportional dilution as certification method and calculate real time particle measurement
DCS is a real-time particle sensor, based on diffusion charge principle
DCS signal is calibrated to actual PM mass post-testing using the filter based method.
The use of the diffusion screens to eliminate PM mass over-reading during DPF regeneration
PM mass measured by filter using a portable exhaust flow rate proportional partial flow dilution system
Real time measurement
Gravimetric Method
This method is approved by EPA alternative approval letter
Filter
Bypass
Cyclone
【Patent No.JP 4956178】

## Slide 15: Method of OBS-PM for NTE
PM mass in Individual NTE events = Total PM × Fraction
NTE-1
NTE-2
NTE-3
NTE-4
NTE-5
Real time measurement
Fraction of NTE
Filter Filter Filter Filter Filter
Bypass Bypass Bypass Bypass Bypass Bypass
Gravimetric method
Total PM mass
DCS signal (mm/cm3)
The ratio of integrated real time particle signal to PM mass (from PM filter) is used to calculate real time PM mass for US NTE and EU moving average window methods
allows subtraction of the PM mass loaded during invalid NTE events < 30 seconds
Also allows determination of real time PM concentration for Moving Average Window

## Slide 16: Real time PM (DCS-100)
Specification:
Measurement range: 0 ... 2500 mm/cm3
Sensitivity limit: 0.01 mm/cm3
Sample flow: 2.5 liters/min. (approx.)
Measures Particle Length d1.13

## Slide 17
Filter change method
Open the door
Step 2
Release Filter lock
Step 3
Remove Filter and holder
Step 1

## Slide 18
Y=1.046x – 0.0042
R2 = 0.9947
OBS-ONE-PM Correlation with MDLT
Correlation between laboratory HORIBA MDLT partial flow and HORIBA OBS-ONE-PM PEMS
Good agreement across wide range of PM emissions

## Slide 19
Proportionality of the PEMS-PM sample flow to engine exhaust flow during FTP cycle
Fast response of flow controller was verified
Y=0.0014X -0.0013
R2=0.9971
Proportionality of PM Sample Flow

## Slide 20
High emission at sloping road
Good linearity at real time testing
Sample flow rate proportionality
Evaluation by on board testing
OBS-ONE-PM Real-Time PM Test Result

## Slide 21
System Features
OBS-ONE-PM (Particulate Mass)
OBS-ONE-PN (Particle Number)

## Slide 22
HORIBA utilized the CPC based system
  Main advantage of CPC
    Particle count principle (R83)
    No particle size dependency
    High accuracy and high sensitivity
  HORIBA’s CPC does not have a liquid reservoir of IPA (soaked wick only)
    “Easy” and “Safe” operation by using Cartridge
HORIBA PN-PEMS Solution
Notes: 22

## Slide 23
System configuration
  Two diluters
  Volatile particle remover: Catalytic Stripper (350 degC)
  CPC (Isopropanol based)
2nd diluter
1st diluter
Catalytic stripper
CPC
OBS-ONE-PN flow schematics
Sample In

## Slide 24: Detection Efficiency
Chart: Particle Diameter (nm)
| 15 | 20 | 23 | 41 | 50 | 60 | 70 | 85 | 100 | 200
SATT:29C | 0.046021 | 0.22578 | 0.32911 | 0.73627 | 0.77998 | 0.83139 | 0.87128 | 0.91019 | 0.95012 | 1
SATT:30C | 0.045778 | 0.32615 | 0.44847 | 0.59177 | 0.65987 | 0.73984 | 0.80581 | 0.83719 | 0.86713 | 0.89283 | 0.93235 | 0.96097 | 1
SATT:31C, CONDT:19C | 0.14714 | 0.31625 | 0.37498 | 0.60527 | 0.68853 | 0.74824 | 0.85539 | 0.87764 | 0.90244 | 0.967 | 1.0081 | 1
Detection efficiency of the system
VL-05
Detection efficiencies were within all criteria defined by the draft RDE regulation

## Slide 25: Response Time
VL-07
 | Criteria | Results
Delay time (Td+T10) | ≦5.0 sec | 4.8 Sec
Rise time (T10-T90) | ≦3.5 sec | 2.0 Sec
System Response Time | - | 6.8 Sec
Td+T10
T10-T90
90%
10%

## Slide 26: HORIBA PN-PEMS result (Phase II)
Regeneration (Volatile particle removal)
HORIBA PN-PEMS
Data obtained by JRC (published in the evaluation report
“Feasibility study on the extension of the RDE procedure to PN“)

## Slide 27: HORIBA PN-PEMS result (Phase II)
NEDC (Cold, PMP@CVS)
NEDC (Hot, PMP@tailpipe)
HORIBA
HORIBA
Real time
Data obtained by JRC (published in the evaluation report
“Feasibility study on the extension of the RDE procedure to PN“)