## Slide 1: Challenges for PEMS under RDE Regulations Dr David Booker, Sensors Inc. CECERT PEMS Conference March, 2016

## Slide 2: Introduction
Background to the EU Legal Process
State of the Play with the RDE Program
Challenges for PEMS
  Gaseous
  PN

## Slide 3: Comitology in the EU
Five Key Reasons behind Comitology
Comitology in the EU references a process by which the EU law is modified or adjusted and takes place within "comitology committees" that assist in the making, adoption, and implementation of EU laws. Broadly speaking, before it can implement an EU legal act, the Commission must consult, for the detailed implementing measures it proposes, a committee where every EU country is represented. For example, the Technical Committee Motor Vehicles (TCMV) committee. In addition “Expert Groups” are formed to develop the scope, legal text and “tools” (eg EMROAD / CLEAR).
Regulation (EC) 715/2007

## Slide 4: RDE - TCMV 2015
Client Earth, 2015
On 19th May 2015, the 1st package (testing procedures) and on 28 October 2015, (conformity factors) the TCMV voted / passed the implementing measures to introduce Real Driving Emission (RDE) tests for emissions from light duty vehicles, under the framework of Regulation (EC) No 715/2007.1 with a clear majority 28 EU Member States (only the Netherlands opposed the 2nd package).
Conformity Factor 2.1 (2017) > 1.5 (2020)

## Slide 5: What Happened in the EP
The European Parliament has the power to oppose the implementing measures adopted by the Commission, if they exceed the implementing powers granted by the EU legislature or are not compatible with the aim or the content of the Euro 6 Regulation.
After TCMV voted in October 2015 the Opposition Stated to gain momentum.

## Slide 6: The EP Vote
2nd February 2016
For: 323
Against: 317
Most of the Socialist EU Parliamentarians wanted to reject the Commissions Proposal. If they had succeeded the EC would have had 3 choices:
Go back to TCMV (unlikely)
Legally challenge the EP (very unlikely), or
Implement the regulation in Euro VII
All could have resulted in a 3-5 year delay!

## Slide 7: EC Major Concessions:
European Commission agreed to review the Conformity Factors on an annual basis. Any revisions would be enforceable 4 years later
  Expert Group (First meeting Feb 2016)
European Commissioner speech to Parliament states that aim is to have cars designed to meet a CF of 1.0 by 2023

## Slide 8: What’s Next

## Slide 10
NDUV NOX (NO +NO2 )Error
Question - 5+5 or 10ppm
CLD NOX <5ppm

## Slide 11: PEMS Measurement Uncertainty
ACEA and JRC Analysis for Gaseous Measurements ONLY to date.

## Slide 12: Drift (No drift Correction Allowed)

## Slide 13: Time Alignment

## Slide 14: Exhaust Flow Measurement

## Slide 15: Other Errors

## Slide 16: CF ACEA (1.7), EC 1.6 (2017) > 1.2 (2020). Current Regulation 2.1 > 1.5

## Slide 17: Background
Correlation / Validation testing over WLTC “highly recommended” for Packages 1 & 2 RDE testing
  Limits (% and Absolute) provided in proposed regulations
  Example Correlation testing performed at UTAC on 1.6 liter diesel vehicle

## Slide 18: Correlation to Laboratory – WLTC Cycle 1.6 liter Diesel at UTAC, 10/29/2014
Chart: Exhaust Mass Flow Rate
Chart: 

## Slide 19: Correlation to Laboratory – WLTC Cycle 1.6 liter Diesel at UTAC, 10/29/2014
Chart: CO2 Wet Concentration
Chart: 

## Slide 20: PEMS Testing – RDE, Development, Fuel Economy
Requirement: Errors <5%
A Factor of 2 better than the current RDE validation test criteria.

## Slide 21: Correlation to Laboratory – WLTC Cycle 1.6 liter Diesel at UTAC, 10/29/2014
Chart: CO Wet Concentration

## Slide 22: Correlation to Laboratory – WLTC Cycle 1.6 liter Diesel at UTAC, 10/29/2014
Chart: NOx, ppm
Chart: NOx, ppm

## Slide 23: Correlation to Laboratory – WLTC Cycle 1.6 liter Diesel at UTAC, 10/29/2014
Chart: NO, ppm
Chart: NO2, ppm

## Slide 24: Correlation Evaluation: PEMs vs CVS Bag
Chart: Percent Difference vs Tolerance
| CO2 mass | CO mass | NOx mass
% Difference (Abs) | 0.00051384 | 0.18846 | 0.041103
Tolerance | 0.1 | 0.15 | 0.15
Chart: Absolute Difference vs Tolerance
| CO2 mass | CO mass | NOx mass
Difference (Abs) | 0.082677 | 11.1193 | 29.9227
Tolerance | 10 | 150 | 15
Pass
Pass
Pass
All gases pass requirements. Low CO levels result in higher %error, but well within the .15 g/km absolute tolerance.

## Slide 25: SEMTECH CPN compliant to the PMP Standard - UNECE Reg. 83
SENSORS CPC Module

## Slide 26: Ref. Technology (Tail-pipe to CVS)

## Slide 27: VPR-Catalyst
Catalyst from Catalytic Instruments
  PMP Specifications
  67% penetration at 10 nm
PND1 Heated Dilution
10-20:1
Heated Catalyst
300 °C
Picture - Feasibility study on the extension of
the Real Driving Emissions (RDE)
procedure to Particle Number (PN), JRC 2014

## Slide 28: VPR-Catalyst – Experimental Verification
Catalyst from Catalytic Instruments
  PMP Specifications
  67% penetration at 10 nm
Stage 1 Heated Dilution
10-20:1
Heated Catalyst
300 °C
Source - Feasibility study on the extension of
the Real Driving Emissions (RDE)
procedure to Particle Number (PN), JRC 2014

## Slide 29
Aerosol Laboratory Evaluation
Linearity of CPC
Size Classification
No Loss correction
applied during Study
D50 ca 23 nm,
Correlation with Diluter

## Slide 30: SEMTECH CPN vs TSI SPC
Data obtained at JRC ISPRA
CPC Evaluation:
Exhaust of CPN module measured simultaneously by TSI CPC and Sensors CPC during chassis dyno testing
Note LOG Axis
(Highlights dynamic range and excellent correlation)

## Slide 31
Test Cell Comparison
SEMTECH CPN and Horiba PMP
Data Supplied by
..Source Confidential......
The charts are normalized
Cycle averages within 20% to reference PMP

## Slide 32
Test Cell Comparison
SEMTECH CPN and AVL 489 PMP
Data Supplied by JRC Dr.Francesco Riccobono
Cycle averages within <20% to reference PMP

## Slide 33
Heavy Duty Test Cell Comparison
SEMTECH CPN and AVL 489 PMP
Data Supplied by Cummins Dr. Shirish Shimpi
Demonstrated excellent Linearity and repeatability

## Slide 34
Thank You For Your Attention