## Slide 1: Improving Flow Rate Exhaust Measurement when utilizing “Small” PEMS
John Karim
In-Use Retrofit Section
Mobile Source Control Division
California Air Resources Board
April 11,2013

## Slide 2
Objective
Evaluate the exhaust emissions of ON/OFF road Motorcycles under “real world” conditions

## Slide 3
Challenges

## Slide 4: Requirements
Light Weight:
  No significant change to center of gravity of the MC
  No effect on MC’s maneuverability
  Compact
  Weather resistant
  Safety concerns (FID fuel)
Measures THC, CO, CO2, NOx
  PM
  NH3
Power supply:
  Own power (for short periods)
  12V- 15A max (~5A)
Works on all types of engines:
  Carburetors/Fuel Injection
  2/4 stroke
  Diesel, Gasoline, etc…
Exhaust flow rate measurement
  ECU
  No ECU
  Minimize pulsation effect

## Slide 5: Small PEMS (AXION R/S System)
NH3 Module
NDIR
Continuous Gaseous Measurement (sec-by-sec) of CO2, CO, NOX, THC, PM, and NH3 emissions
CO2, CO, and THC by NDIR
NOX by Electrochemical Cell
O2 by Electrochemical Cell
PM by Laser Light Scattering
NH3 by Tunable Diode Laser Spectrometry (TDLS)
NOx Sensor
O2 Sensor
PM Sensor

## Slide 6: Small PEMS (AXION R/S System) Cont.
Other Continuous Measurement Components
Vehicle ECU Communication by OBD LDScanner and J1708/J1939 Scanner.
Intake Air Temperature (IAT) by Thermistor
Engine Speed (RPM) by Piezoelectric Tachometer, Optical Tachometer, or Inductive Tachometer
Manifold Absolute Pressure (MAP) by MAP Transducer
Global Positioning System (GPS) provide latitude, longitude, vehicle speed, and bearing.
Ambient Condition by Weather Envoy

## Slide 7: Specifications of Small PEMS
Gas | Measurement Range | Accuracy | Repeatability | Noise | Resolution
CO2 | 0 to 16 % | 3% rel or 0.3% abs | 2% rel or 0.1% abs | 0.8% rel or 0.1% abs | 0.01 vol %
CO | 0 to 10 % | 3% rel or 0.02% abs | 2% rel or 0.02% abs | 0.8% rel or 0.01% abs | 0.001 vol %
THC | 0 to 10,000 ppm | 3% rel or 4 ppm | 2% rel or 3 ppm | 0.8% rel or 2 ppm | 1 ppm
NOX | 0 to 4,000 ppm | 4% rel or 25 ppm | 3% rel or 20 ppm | 1% rel or 10 ppm | 1 ppm
PM | 0 to 300 mg/m3 | NA | NA | NA | 0.01 mg/m3
O2 | 0 to 25% | 3% rel or 0.1% abs | 3% rel or 0.1% abs | 1.5% rel or 0.1% abs | 0.01 vol %
NH3 | 0 to 500 ppm | 2% rel or 2 ppm | NA | 0.9 ppm | 0.1 ppm

## Slide 8: Continuous Measurement
NDIR
NOx Sensor
O2 Sensor
System 1
System 2
NDIR
NOx Sensor
O2 Sensor
Notes: Dual system to enable continuous measurement during zero checks which can be programed for a user-difined. interval

## Slide 9: Linearity Check
Gas Analyzers #1
Gas Analyzers #2

## Slide 10: Gas Analyzer #1 and #2 (Linearity Check)
16
10
4,000
10,000
Range
Instrument noise levels

## Slide 11: PEMS Testing on SS Dyno (cell6)
Use Small PEMS (Axion-R/S system) and SEMTECH-ECOSTAR
Test 3 motorcycles in dyno lab with four-steady state cycle (idle, 15 mph, 30 mph, and 50 mph)
Check the real time concentration profiles and the emissions
CVS Tunnel
CVS Instrument Train
Bag #1
Bag #2
Bag #3
Dilution Air
SEMTECH-EFM
Exhaust
AXION R/S
SEMTECH-ECOSTAR

## Slide 12: Concentration Profile Motorcycle 2008 Harley Davidson FXDWG 1584 c.c.
. Axion R/S
. SEMTECH-ECHOSTAR

## Slide 13: Concentration Profile Motorcycle 2006 Honda VTX 1300C 1312 c.c.
. Axion R/S
. SEMTECH-ECHOSTAR

## Slide 14: Concentration Profile Motorcycle 2008 Yamaha YZFR1XCR 1000 c.c.

## Slide 15: Summary of Motorcycle Steady State Tests (Total Mass)

## Slide 16: PEMS Testing on Gasoline Vehicle
Use Small PEMS (Axion-R/S system) and SEMTECH-AEA (including SEMTECH-DS)
Test in dyno lab with FTP-75 Cycle
Check the real time concentration profiles and the emissions

## Slide 17: Concentration Profile Gasoline Vehicle 2008 Chevrolet Impala 3.5L Test: FTP-75 Cycle With Cold Start

## Slide 18: Concentration Profile Gasoline Vehicle 2008 Chevrolet Impala 3.5L Test: FTP-75 Cycle With Hot Start

## Slide 19: Summary of the Gasoline Vehicle Tests (Total Mass) 2008 Chevrolet Impala 3.5L FTP-75 Cycle

## Slide 20: Volumetric Efficiency Factor
Volumetric Efficiency refers to the efficiency with which the engine can move the charge into and out of the cylinders.
Exhaust flow rate is dependent on IAT, MAP, and RPM

## Slide 21: Volumetric Efficiency Determination (Dyno 6)
Estimate the Volumetric Efficiency Factor (VEF) based on real time results from “dyno modal data” and “Axion R/S system”
Check the real time mass concentration profiles and the emissions
CVS Tunnel
CVS Instrument Train
Bag #1
Bag #2
Bag #3
Dilution Air
SEMTECH-EFM
Exhaust
AXION R/S
SEMTECH-AEA
Use Small PEMS (Axion-R/S system)
Test in Dyno 6 with FTP-72 Cycle

## Slide 22: List of Motorcycles Tested for VE Correction
No | Model | Model Year | Engine Displacement
01 | Yamaha YZFR1 XCR | 2008 | 998 cc
02 | Kawasaki KLR650 | 2008 | 651 cc
03 | Suzuki SV650S | 2006 | 645 cc
04 | Honda VTX1300C | 2006 | 1312 cc
05 | Harley Davidson FXDWG | 2008 | 1584 cc
06 | Suzuki DRZ400SM | 2008 | 398 cc
07 | Harley Davidson Softail | 1995 | 1340 cc
08 | Honda CBR600RR | 2008 | 600 cc
09 | Honda CMX250C | 2006 | 234 cc
10 | Vespa GTS250 | 2007 | 244 cc
11 | Yamaha Vino 125 | 2007 | 125 cc

## Slide 23: 2008 Suzuki 398 cc

## Slide 24: 2008 Suzuki 398 cc

## Slide 25: 2008 Suzuki 398 cc

## Slide 26: 2008 Suzuki 398 cc

## Slide 27: 2008 Suzuki 398 cc

## Slide 28: Honda 1312 cc

## Slide 29: Kawasaki 651 cc

## Slide 30: Suzuki 645 cc

## Slide 31: Harley Davidson 1584 cc

## Slide 32: Suzuki 398 cc
Test 1
Test 2

## Slide 33: Yamaha 998 cc

## Slide 34: Post-PEMS Testing Confirmation (Cell 5)
Use Small PEMS (Axion-R/S system)
Test in dyno lab with FTP-75 Cycle
Evaluate the established volumetric efficiency (VEF) tables on the re-test in dyno lab
Check the real time mass concentration profiles and the emissions
CVS Tunnel
CVS Instrument Train
Bag #1
Bag #2
Bag #3
Dilution Air
Exhaust
AXION R/S

## Slide 35: Harley Davidson 1584 cc

## Slide 36: 2008 Harley Davidson FXDWG 1584 cc (Test #1)

## Slide 37: 2008 Harley Davidson FXDWG 1584 cc (Test #2)

## Slide 38: Summary and Conclusion
There is a need for a “small” PEMS
  Small
  Light weight
  Easy hook-up
  One person installation and operation
  Low Energy Usage
  Dual analyzers
Further enhancements are needed
  Sample conditioning
  Dual range analyzers (even for 1065 PEMS)
  Better techniques to calculate VE (IAT, MAP, and RPM)
  Better technology for HC
  Include corrected NOx
  Provide Post Processing software
Built from ground up
Adequate measurement can be achieved

## Slide 39: Thank you
In-Use Retrofit Section
Dr. Jesse Tu, Stephanie Maalouf, George Gatt, Gloria Pak, Tony Nassar, Svetlana Espinosa, Tung Tran, Robert Cabral, Anne Mackhampio, Patrick Chang, Nick Le
Special Thanks to
Mr. Brian Beckmann
Mechanical and Aerospace Engineer
GlobalMRV Inc.
Sharon Lemieux, Chief
Heavy-Duty Diesel In-Use Strategies Branch