## Slide 1: SEMTECH® LASAR Trace Gas Analyzer
SEMTECH® LASAR:
High Speed Tools for The In-use Measurements of Nitrogen-Containing Species.
Bertrand S. Lanher, PhD.
Sensors, Inc.
www.sensors-inc.com
2012 PEMS Conference and Workshop – UCR/CE-CERT

## Slide 2: SEMTECH® LASAR Trace Gas Analyzer
SEMTECH® LASAR:
Improving upon TDLAS for
ppb-level, interference-free
in-use measurements of nitrogen-containing (and other) species.
2012 PEMS Conference and Workshop – UCR/CE-CERT

## Slide 3: SEMTECH® LASAR Trace Gas Analyzer
Part 1 - Spectroscopy

## Slide 4: SEMTECH® LASAR Trace Gas Analyzer
STEP 1
Replace the Herriot (multipass) gas cell with a Fabry-Perot ring resonator.
Advantage:
Increase pathlength from a few meters to up to 20 km.
Benefit:
Increase signal (and sensitivity) by 1,000x

## Slide 5: SEMTECH® LASAR Trace Gas Analyzer
STEP 1
Replace the Herriot (multipass) gas cell with a Fabry-Perot ring resonator.
Advantage:
Increase pathlength from a few meters to up to 20 km.
Benefit:
Increase signal (and sensitivity) by 1,000x
TDL’s used with hyper-reflective gas cells outperform QCL with multipass gas cells.

## Slide 6: SEMTECH® LASAR Trace Gas Analyzer
STEP 2
Reduce laser light source emission bandwidth from
1 MHz to 10 KHz
by means of optical feedback.
Advantage:
Improves laser light emission purity.
Benefits:
Reduces noise (and improves sensitivity) by 100x.
Improves instrument response linearity.
0.1 pm WL accuracy.
0.005 cm-1 optical resolution
Feedback
Mode Locking
No Feedback
No Mode Locking

## Slide 7: SEMTECH® LASAR Trace Gas Analyzer
STEP 3
Apply current ramp to laser light source to change very quickly emission wavelength.
Advantage:
Collects broad spectrum information.
Benefits:
Measures more than 1 gas with a single laser light source.
Zero absorbance region provide auto-zero of the analyzer.
l1
Laser
Intensity
(mA)
20
80
l (nm)
lT + dlI
lT
t(s)
l2
l3
l199
l200
100 ms
0.5 ms

## Slide 8: SEMTECH® LASAR Trace Gas Analyzer
STEP 4
Apply phase modulation to the laser light emission.
Advantage:
Widens spectral range.
Benefit:
Measure more than 1 gas with a single laser light source.
Zero absorbance region provide auto-zero of the analyzer.

## Slide 9: SEMTECH® LASAR Trace Gas Analyzer
STEP 5
Real-time calibration against primary, traceable standards
Advantage:
Primary measurement method.
Benefit:
Improves accuracy.
Full traceability of data/results.
REFERENCE (HiTran, GEISA, etc.)
EXPERIMENTAL (Direct Absorbance)
Center Wavelength
(identification)
Lorenz Coefficient
(symmetry for ID & conc.)
Gauss Coefficient
(Concentration)
Center Wavelength
(identification)
Lorenz Coefficient
(symmetry for ID & conc.)
Gauss Coefficient
(Concentration)
GAS
concentration
FIT

## Slide 10: SEMTECH® LASAR Trace Gas Analyzer
SPECTROSCOPY – SUMMARY OF IMPROVEMENTS
5 Orders of Magnitude Sensitivity Improvement
0.005 cm-1 Optical Resolution
Up To 400 Wavelengths scanned every 100 milliseconds
Built-in Zero Information (analyzer auto-zeroed @ 10Hz)
Multiple Gases Measured with a Single Laser.
Wider Dynamic Range (> 1:5,000 @ 1 Hz; 1:50,000 @ 0.03 Hz)
Primary Measurement Method (improved accuracy + traceability)

## Slide 11: SEMTECH® LASAR Trace Gas Analyzer
Part 2 - Sampling

## Slide 12: SEMTECH® LASAR Trace Gas Analyzer
Low Pressure Sampling
Operate the sampling line and gas analyzer between 50-100 mbar absolute.
Advantage 1:
Lowers dew point.
Benefit 1:
Operates sampling lines and analyzer at room temperature.
H2O(g) % vol.
Dew Point T°C
@ 50 mbar
10 % vol.
- 2.73°C (27.1F)
20 % vol.
+ 6.96°C (44.5F)
30 % vol.
+13.01°C (55.4F)

## Slide 13: SEMTECH® LASAR Trace Gas Analyzer
Low Pressure Sampling
Operate the sampling line and gas analyzer between 50-100 mbar absolute.
Advantage 2:
Operates at low intake flow rate (<0.33 L/min), yet improves sample transfer speed.
Benefit 2:
Reduces overall analyzer response time.
Sample Transfer Speed between 1 and 4 m/sec. when using ¼” tubing (4/6 mm) and for flow rates of 5 to 20 L/hour, respectively.

## Slide 14: SEMTECH® LASAR Trace Gas Analyzer
Low Pressure Sampling
Operate the sampling line and gas analyzer between 50-100 mbar absolute.
Advantage 3:
Very low intake flow rate (<0.33 L/min) reduces fouling as a function of time.
Benefit 3:
Reduces fouling and optics maintenance.

## Slide 15: SEMTECH® LASAR Trace Gas Analyzer
Low Pressure Sampling
Operate the sampling line and gas analyzer between 50-100 mbar absolute.
Advantage 4:
Allows low temperature operation and reduces collision rate, therefore reducing chemical reaction rate in the sampling line.
Benefit 4:
Improved accuracy.
Sample
sample
[NH3] = x
Chemical Reactions
[NH3] = y (>< x)
k(T) = Zρ.eEa/RT
Reaction Coef. Rate
a.k.a. Rate Constant
Collision frequency
Reaction activation energy
Temperature
Gas constant
Steric factor
Z = NAσAB.(8kBT/πµAB)1/2
Collision frequency
Number of molecules / volume
Reaction cross section
Boltzmann’s constant
Temperature
Reduced mass of the
reactants (A and B)

## Slide 16: SEMTECH® LASAR Trace Gas Analyzer
Low Pressure Sampling
Operate the sampling line and gas analyzer between 50-100 mbar absolute.
Advantage 5:
Reduces bandwidth of absorption bands to the point of removing spectral overlapping.
Benefit 5:
Removes interferences.

## Slide 17: SEMTECH® LASAR Trace Gas Analyzer
Low Pressure Sampling
Operate the sampling line and gas analyzer between 50-100 mbar absolute.
Advantage 5:
Reduces bandwidth of absorption bands to the point of removing spectral overlapping.
Benefit 5:
Removes interferences.

## Slide 18: SEMTECH® LASAR Trace Gas Analyzer
SAMPLING – SUMMARY OF IMPROVEMENTS
Room temperature operation (no heated lines required)
Very low flow rate (0.33 L/min.) for low fouling
Faster sample transfer rate (up to 4 m/sec.)
Improved accuracy (lower chemical reaction rate in sampling line)
Removes spectral interferences.

## Slide 19: SEMTECH® LASAR Trace Gas Analyzer
Part 3 - Results

## Slide 20: SEMTECH® LASAR Trace Gas Analyzer
Linearity
0 – 25 ppm N2O
Slope = 1.004;
Intercept = 0.013 ppm
R^2 = 1.000
R^2 = 0.99996
0 – 8,000 ppm H2O(g)

## Slide 21: SEMTECH® LASAR Trace Gas Analyzer
Dynamic Range
0 – 50 ppm N2O calibration
30-sec. Std. Dev.< 1 ppb

## Slide 22: SEMTECH® LASAR Trace Gas Analyzer
Stability
50 ppm span N2O
24-hour ‘zero’ stability
Std. Dev. = 4 ppb
1,000 ppm span NH3.
9-day drift test (no re-zero / no re-span)
Drift = 290 ppb

## Slide 23: SEMTECH® LASAR Trace Gas Analyzer
Accuracy,
Precision,
Reproducibility
± 3 σ noise = ± 10 ppb @ 1 Hz
using 50 ppm N2O span calibration
1 σ repeatability < 1 ppb
1 σ reproducibility < 1 ppb
@ 30 seconds integration using 50 ppm N2O span calibration
+/- 10 ppb
(@ 1 Hz without averaging)

## Slide 24: SEMTECH® LASAR Trace Gas Analyzer
No Interference

## Slide 25: SEMTECH® LASAR Trace Gas Analyzer
Fast Response
N2O LASAR
N2O FTIR
LASAR 10 Hz / 1 Hz 0.33 L/min.
FTIR 5 Hz / 1 Hz 15 L/min.
LASAR 10 Hz / 1 Hz 0.33 L/min.
FTIR 5 Hz / 1 Hz 40 L/min.
NH3 LASAR
NH3 FTIR

## Slide 26: LASAR – Spectroscopy Improvements
Optical Resolution (1)
Wavelength precision (2)
(Laser Purity)
Instrumental Noise (4)
(single scan peak-to-peak)
Pathlength (5)
Multi-component capable (6)
Zero (background) (7)
Instrument Drift (8)
Wavelength Accuracy (3)
(Temperature stabilization)
1.5 GHz
~ 10-5 A.U.
1-25 m
yes
Requires regular re-zero
Yes (requires regular re-calibration)
Drift of 1 pm / °C
1 MHz
~ 10-5 A.U.
1-25 m
no
0.5-8.0 cm-1
0.01 cm-1
0.01 cm-1
1 MHz
~ 10-6 A.U.
1-10 km
no
10 kHz
4.10-10 A.U.
1-20 km
yes
Auto-zero
none
± 0.01 pm
0.005 cm-1
LASAR
FTIR
TDL
CRDS
(1)
Distance between 2 consecutive data points
(2)
Laser bandwidth (width at mid-height)
(3)
Laser temperature regulation dependent
(4)
Single scan peak-to-peak noise (± 3σ)
(5)
Practical pathlength; affects sensitivity
(6)
Using a single IR source / Laser
(7)
Re-zero requirement using zero gas
Drift of zero and span as a function of time

## Slide 27: LASAR – Sampling Improvements
Spectral overlapping (2)
Yes in most cases.
Use Multivariate Chemometrics to model
Intake flow rate (4)
Up to 60 l/min.
> 1 l/min
Sample transfer speed (5)
Dependent on line ID @ flow rate
Sample conditioning (1)
Required to avoid condensable
Remove H2O(g) or heat sampling line
FTIR
TDL
CRDS
(1)
Potential for formation of condensates
(2)
Potential for cross response
(3)
Dead volume affecting true response rate
(4)
Actual sample intake flow rate
(5)
Actual sample transfer speed (overall response time)
(6)
Actual T10-90 based on reported flow rates
& gas cell dead volume.
(7)
Hourly fouling rate based on 1000 mg/m3 particulate
matter concentration and reported flow rates
Gas cell volume (3)
> 250 cm3
> 250 cm3
25-50 cm3
T10-90 (6)
< 1sec. @ 60 l/min.
<3 sec.
Fouling rate (7)
3.6 g/hr
60 mg/hr
none
< 0.33 L/min
1-4 m.sec-1
< 1 sec. @ 20 L/hr
2.5 sec. for NH3
Negligible g/hr.
< 25 cm3
none
LASAR

## Slide 28
SEMTECH® LASAR Trace Gas Analyzer
Released (for ICE applications):
NH3 ( + CO2 + H20 )
N2O ( + NH3 + CH4 )
NO ( + CO2 + H2O )
NO2 ( + CO2 + H2 )
H2 ( + CH4 + CO2 )
H2S ( + CO2 + CH4 )
HCN ( + CO2 + H2O )
CO
O2
SO2
SO3
Custom configurations available.
CONFIGURATIONS

## Slide 29: LASAR – Conclusion

## Slide 30: Contact Information
Sensors, Inc.
6812 State Road
Saline, MI 48176
Phone: 734-429-2100
Fax: 734-429-4080
Sensors Europe GmbH
Feldheider Str. 60
40699 Erkrath, Germany
Phone: +49 (0) 2104-14188-0
Fax: +49 (0) 2104-14188-14
www.Sensors-Inc.com
On Board Emissions Analyzers
Test Cell Emissions Analyzers
I / M Emissions Analyzers
Emission Testing Services
Environmental Applications